## Sheet1
DEFIBRILLATOR
SR # | SPECIFICATIONS
1 | Biphasic transthoracic (external) defibrillator with LCD colour display Synchronized output with ECG.
2 | Energy selection & delivery on control panel and paddles for external defibrillation.
3 | Energy selection and delivery on control panel for internal defibrillation.
4 | Charging Indicator
5 | The energy range should be adjustable for peads and adults up to 200Joules.
6 | Charging Time for full energy should be less than 07 sec
7 | Screen Size of approx. 7inch colored.
8 | Display of HR, ECG through paddles and Lead I.II & III patient cable.
9 | Built in recorder for printing of full summery on standard 80mm paper.
10 | Alarms for High and low Heart rate, low battery warning.
11 | Built-in Rechargeable battery with charger for minimum 100 shocks at max energy.
12 | Auto tester/self check.
13 | External Paddles (Adult, Paed)
14 | AED facility with cable. AED Disposable pad with capability of CPR Feedback on defibrilaltor display.
15 | Pacing facility
16 | AC 220V / 50Hz operated.
17 | Accessories:
18 | Complete with standard accessories, including reusable type Adult, Paediatric.
19 | Local made trolley/cart
20 | Optional (If any):
21 | AED Reusable Pad.
22 | Internal Paddle with switch button for (Adult, Paed, Neonate)
24 | ETCo2
25 | NIBP
26 | Spo2
27 | Disposable pacing pads