## Sheet1
ICU MONITOR (MODULAR)
SPECIFICATIONS
Operating Features and Characteristics:
Non fade TFT/LCD color touch screen display
Electro-surgical interference suppression/protection
Defibrillator protection
Waveform traces speed; 25 / 50mm/sec.
Screen size: min. 15” TFT/LCD colour Touch screen display.
Parameters:
ECG : 12-lead ECG analysis with 10-leadwires.
Numeric: heart rate.
NON-INVASIVE BLOOD PRESSURE (NIBP):
Method: oscillometric principle
Numeric: systolic, diastolic and mean pressure
Selectable auto inflate interval settings
Rising cuff/continuous pressure display.
TEMPERATURE:
Numeric: temperature selectable in ºC/ºF.
PULSE OXIMETRY:
Numeric: 30-100% oxygen saturation measuring range.
Waveform-plethysmograph pulse with pulse strength indication.
ARRHYTHMIA ANALYSIS:
Arrhythmia analysis and st analysis. More than 20 arrhythmia analysis should be availabe in the system
RESPIRATION:
Breath rate display and settable apnea alarms.
Sweep speed; 6.25, 12.5 mm/sec.
OTHER FEATURES:
Trend data; graphical and tabular for 96-hours
ALARMS:
High & low (settable) on all parameters
Visual and audible indication of alarms.
OPTIONAL:
IBP Dual Channel /Three Channel (Procuring agency to specify)
Capnography (EtC02)
Printer Two / Three Channel (Procuring agency to specify
EEG Module
Cardiac Out put
OPERATING REQUIREMENTS :
Ac 220v/50HZ
Built-in /external battery for at least 1.5 - 2 hour ac power failure at full parameter.
Accessories: The system must be complete with all sensors, probes, cables or any other accessories required for measuring all the above selected parameters.
Optional (If any): Qty of Reusable sensors Mounting stand preferably imported or high quality (S.S 304L) Local with lockable draws(procuring agency to choose