## Sheet1
Clinical Specialty | CARDIOLOGY
Generic Name | MODULAR BED SIDE MONITOR
Clinical Purpose | In medicine, monitoring is the observation of a disease, condition or one or several medical parameters over time. It can be performed by continuously measuring certain parameters by using a medical monitor (for example, by continuously measuring vital signs by a bedside monitor), and/or by repeatedly performing medical tests (such as blood glucose monitoring with a glucose meter in people with diabetes mellitus).
Modular bedside monitor for Adult / Neonates/ Peads. The monitor should take different modules for display of vital sign monitor of Adult /Neonate/Peads. (The procuring agency will specify the requirement)
Operating Features and Characteristics:
Non fade TFT,LCD color display
Electro-surgical interference suppression/protection
Defibrillator protection
Freeze and cascade facility.
Screen size: min. 19” TFT, LCD color Touch screen display.
Capability to interface with LAN/WLAN for data transfer
Parameters in module form:
ECG:
Numeric: heart rate.
Waveform : Six Wave forms minimum, real time and freeze ECG trace
NON-INVASIVE BLOOD PRESSURE (NIBP):
Method: Oscillometric principle
Numeric: systolic, diastolic and mean pressure
Selectable auto inflate interval settings
Rising cuff/continuous pressure display.
TEMPERATURE:
Numeric: temperature selectable in ºC/ºF.
PULSE OXIMETRY:
Numeric: 30-100% oxygen saturation measuring range.
Waveform-plethysmograph pulse
Reusable sensor electrode.
ARRHYTHMIA ANALYSIS:
Arrhythmia analysis and st analysis.
RESPIRATION:
Breath rate display and settable apnea alarms.
Sweep speed; 6.25, 12.5 mm/sec.
OTHER FEATURES:
Trend data; graphical and tabular of 96 hours
ALARMS:
High & low (settable) on all parameters
Visual and audible indication of alarms.
OPERATING REQUIREMENTS :
Ac 220v/50HZ
Built-in / External Backup for at least 1 hour ac power failure at full parameter.
Accessories: The system must be complete with all sensors, probes, cables or any other accessories required for measuring all the above selected parameters for neonates / peads and Adults.
Optional (If any):
(Procuring agency will select according to its requirement)
IBP Two/ Three/four Channel module , Capnography (EtC02) module
Cardiac Output Module , EEG Module
Printer Two / Three Channel
Qty of Reusable sensors
Mounting stand preferably imported or high quality (S.S 304L) Local with lockable